# Scintillation Detectors for MUSE

# Design Requirements

- Average detection efficiency ε > 0.95.

- Average time resolution, σ < 70 ps.

# Quality Assurance

- BC-404 Scintillators are inspected for damages, inclusions, and refraction index inhomogeneities

- Hamamatsu R9779 PMTs are tested for signal integrity, signal-to-noise ratio, gain, HV requirements, and magnetic field shielding

- Counter Pre-Check: Counters are inspected for void-free glued PMT-to-scintillator transition and light tightness, set to final gain-balanced HVs

- Counter Full-Check: With Three-Bar-Method position dependent and overall time resolutions, effective speed of light, left and right attenuation lengths (BAL and TAL) are programmatically analyzed and automatically stored

- Database: All acquired information is stored and retraceable